## Transmission
TL10X-2P Transmission
|  | Wavelength (nm) | % Transmission
|  | 400 | 68.451
|  | 405.03 | 72.77
|  | 410.05 | 74.766
|  | 415.08 | 75.642
Product Raw Data |  | 420.1 | 76.211
Thorlabs Super Apochromatic Microscope Objective, 10X Magnification |  | 425.13 | 76.413
|  | 430.15 | 76.89
Item # | TL10X-2P | 435.18 | 77.598
DISCLAIMER: The data presented here are typical. Slight variations in performance data will occur from lot to lot. Please contact Technical Support with any questions regarding the use or reliability of this data. |  | 440.2 | 78.397
|  | 445.23 | 79.144
|  | 450.25 | 79.725
|  | 455.28 | 80.081
|  | 460.3 | 80.187
|  | 465.33 | 79.988
This data may be used in publications. However, please cite Thorlabs as the source. |  | 470.35 | 79.643
|  | 475.38 | 79.232
|  | 480.4 | 78.83
Additional Information: |  | 485.43 | 78.49
|  | 490.45 | 78.28
|  | 495.48 | 78.233
|  | 500.5 | 78.354
|  | 505.53 | 78.558
|  | 510.55 | 78.917
|  | 515.58 | 79.404
|  | 520.6 | 79.988
|  | 525.63 | 80.631
|  | 530.65 | 81.297
|  | 535.68 | 81.948
|  | 540.7 | 82.552
|  | 545.73 | 83.082
|  | 550.75 | 83.486
|  | 555.78 | 83.673
|  | 560.8 | 83.743
|  | 565.83 | 83.695
|  | 570.85 | 83.535
|  | 575.88 | 83.276
|  | 580.9 | 82.932
|  | 585.93 | 82.514
|  | 590.95 | 82.048
|  | 595.98 | 81.553
|  | 601.01 | 81.052
|  | 606.03 | 80.573
|  | 611.06 | 80.115
|  | 616.08 | 79.692
|  | 621.11 | 79.313
|  | 626.13 | 78.99
|  | 631.16 | 78.724
|  | 636.18 | 78.517
|  | 641.21 | 78.372
|  | 646.23 | 78.288
|  | 651.26 | 78.277
|  | 656.28 | 78.365
|  | 661.31 | 78.508
|  | 666.33 | 78.7
|  | 671.36 | 78.938
|  | 676.38 | 79.216
|  | 681.41 | 79.528
|  | 686.43 | 79.869
|  | 691.46 | 80.231
|  | 696.48 | 80.61
|  | 701.51 | 80.992
|  | 706.53 | 81.363
|  | 711.56 | 81.732
|  | 716.58 | 82.093
|  | 721.61 | 82.441
|  | 726.63 | 82.772
|  | 731.66 | 83.081
|  | 736.68 | 83.366
|  | 741.71 | 83.623
|  | 746.73 | 83.851
|  | 751.76 | 84.046
|  | 756.78 | 84.208
|  | 761.81 | 84.337
|  | 766.83 | 84.432
|  | 771.86 | 84.492
|  | 776.88 | 84.521
|  | 781.91 | 84.517
|  | 786.93 | 84.482
|  | 791.96 | 84.42
|  | 796.98 | 84.33
|  | 802.01 | 84.204
|  | 807.04 | 84.039
|  | 812.06 | 83.855
|  | 817.09 | 83.654
|  | 822.11 | 83.44
|  | 827.14 | 83.214
|  | 832.16 | 82.98
|  | 837.19 | 82.739
|  | 842.21 | 82.494
|  | 847.24 | 82.247
|  | 852.26 | 82
|  | 857.29 | 81.755
|  | 862.31 | 81.515
|  | 867.34 | 81.28
|  | 872.36 | 81.052
|  | 877.39 | 80.834
|  | 882.41 | 80.625
|  | 887.44 | 80.427
|  | 892.46 | 80.242
|  | 897.49 | 80.07
|  | 902.51 | 79.917
|  | 907.54 | 79.784
|  | 912.56 | 79.667
|  | 917.59 | 79.564
|  | 922.61 | 79.478
|  | 927.64 | 79.407
|  | 932.66 | 79.353
|  | 937.69 | 79.315
|  | 942.71 | 79.293
|  | 947.74 | 79.288
|  | 952.76 | 79.298
|  | 957.79 | 79.324
|  | 962.81 | 79.366
|  | 967.84 | 79.423
|  | 972.86 | 79.494
|  | 977.89 | 79.58
|  | 982.91 | 79.679
|  | 987.94 | 79.792
|  | 992.96 | 79.917
|  | 997.99 | 80.055
|  | 1003.02 | 80.208
|  | 1008.04 | 80.376
|  | 1013.07 | 80.553
|  | 1018.09 | 80.739
|  | 1023.12 | 80.934
|  | 1028.14 | 81.137
|  | 1033.17 | 81.346
|  | 1038.19 | 81.561
|  | 1043.22 | 81.782
|  | 1048.24 | 82.007
|  | 1053.27 | 82.235
|  | 1058.29 | 82.467
|  | 1063.32 | 82.693
|  | 1068.34 | 82.916
|  | 1073.37 | 83.14
|  | 1078.39 | 83.362
|  | 1083.42 | 83.584
|  | 1088.44 | 83.803
|  | 1093.47 | 84.019
|  | 1098.49 | 84.231
|  | 1103.52 | 84.439
|  | 1108.54 | 84.641
|  | 1113.57 | 84.837
|  | 1118.59 | 85.026
|  | 1123.62 | 85.207
|  | 1128.64 | 85.38
|  | 1133.67 | 85.545
|  | 1138.69 | 85.699
|  | 1143.72 | 85.843
|  | 1148.74 | 85.977
|  | 1153.77 | 86.098
|  | 1158.79 | 86.208
|  | 1163.82 | 86.306
|  | 1168.84 | 86.391
|  | 1173.87 | 86.462
|  | 1178.89 | 86.519
|  | 1183.92 | 86.562
|  | 1188.94 | 86.591
|  | 1193.97 | 86.605
|  | 1199 | 86.603
|  | 1204.02 | 86.571
|  | 1209.04 | 86.519
|  | 1214.07 | 86.451
|  | 1219.1 | 86.368
|  | 1224.12 | 86.269
|  | 1229.15 | 86.153
|  | 1234.17 | 86.022
|  | 1239.2 | 85.875
|  | 1244.22 | 85.711
|  | 1249.25 | 85.532
|  | 1254.27 | 85.336
|  | 1259.3 | 85.125
|  | 1264.32 | 84.898
|  | 1269.35 | 84.656
|  | 1274.37 | 84.398
|  | 1279.4 | 84.125
|  | 1284.42 | 83.837
|  | 1289.45 | 83.535
|  | 1294.47 | 83.218
|  | 1299.5 | 82.887
|  | 1304.52 | 82.543
|  | 1309.55 | 82.184
|  | 1314.57 | 81.813
|  | 1319.6 | 81.429
|  | 1324.62 | 81.033
|  | 1329.65 | 80.624
|  | 1334.67 | 80.204
|  | 1339.7 | 79.773
|  | 1344.72 | 79.33
|  | 1349.75 | 78.878
|  | 1354.77 | 78.415
|  | 1359.8 | 77.943
|  | 1364.82 | 77.461
|  | 1369.85 | 76.971
|  | 1374.87 | 76.472
|  | 1379.9 | 75.965
|  | 1384.93 | 75.451
|  | 1389.95 | 74.93
|  | 1394.98 | 74.403
|  | 1400 | 73.869
|  | 1400 | 70.8

## Axial Color Correction
TL10X-2P Axial Color Correction
|  | Wavelength (nm) | Theoretical Axial Focal Shift (mm)
|  | 400 | -2.3018e-05
|  | 407.6923 | 0.00017323
|  | 415.3846 | 0.00030202
|  | 423.0769 | 0.00038589
Product Raw Data |  | 430.7692 | 0.00044068
Thorlabs Super Apochromatic Microscope Objective, 10X Magnification |  | 438.4615 | 0.00047742
|  | 446.1538 | 0.0005037
Item # | TL10X-2P | 453.8462 | 0.00052468
DISCLAIMER: The data presented here are typical. Slight variations in performance data will occur from lot to lot. Please contact Technical Support with any questions regarding the use or reliability of this data. |  | 461.5385 | 0.00054372
|  | 469.2308 | 0.00056296
|  | 476.9231 | 0.00058362
|  | 484.6154 | 0.00060631
|  | 492.3077 | 0.0006312
|  | 500 | 0.00065815
This data may be used in publications. However, please cite Thorlabs as the source. |  | 507.6923 | 0.00068684
|  | 515.3846 | 0.00071681
|  | 523.0769 | 0.00074754
Additional Information: |  | 530.7692 | 0.00077848
|  | 538.4615 | 0.00080904
|  | 546.1538 | 0.00083869
|  | 553.8462 | 0.00086687
|  | 561.5385 | 0.00089308
|  | 569.2308 | 0.00091684
|  | 576.9231 | 0.00093771
|  | 584.6154 | 0.0009553
|  | 592.3077 | 0.00096923
|  | 600 | 0.00097918
|  | 607.6923 | 0.00098484
|  | 615.3846 | 0.00098595
|  | 623.0769 | 0.00098227
|  | 630.7692 | 0.0009736
|  | 638.4615 | 0.00095975
|  | 646.1538 | 0.00094055
|  | 653.8462 | 0.00091588
|  | 661.5385 | 0.00088559
|  | 669.2308 | 0.00084961
|  | 676.9231 | 0.00080783
|  | 684.6154 | 0.00076018
|  | 692.3077 | 0.0007066
|  | 700 | 0.00064703

## Strehl Ratio
TL10X-2P Strehl Ratio
|  | Measured Field (°) | Measured Strehl Ratio | Theoretical Field (°) | Theoretical Strehl Ratio
|  | -3.25 | 0.867 | -3.575 | 0.97761
|  | -1.625 | 0.915 | -3.432 | 0.98058
|  | 0 | 0.969 | -3.289 | 0.98134
|  | 1.625 | 0.915 | -3.146 | 0.98056
Product Raw Data |  | 3.25 | 0.826 | -3.003 | 0.98156
Thorlabs Super Apochromatic Microscope Objective, 10X Magnification |  |  |  | -2.86 | 0.9815
|  |  |  | -2.717 | 0.98061
Item # | TL10X-2P |  |  | -2.574 | 0.97914
DISCLAIMER: The data presented here are typical. Slight variations in performance data will occur from lot to lot. Please contact Technical Support with any questions regarding the use or reliability of this data. |  |  |  | -2.431 | 0.9773
|  |  |  | -2.288 | 0.97521
|  |  |  | -2.145 | 0.97295
|  |  |  | -2.002 | 0.97058
|  |  |  | -1.859 | 0.96811
|  |  |  | -1.716 | 0.96552
This data may be used in publications. However, please cite Thorlabs as the source. |  |  |  | -1.573 | 0.9628
|  |  |  | -1.43 | 0.95991
|  |  |  | -1.287 | 0.95681
Additional Information: |  |  |  | -1.144 | 0.95484
|  |  |  | -1.001 | 0.95369
|  |  |  | -0.858 | 0.95238
|  |  |  | -0.715 | 0.95086
|  |  |  | -0.572 | 0.94907
|  |  |  | -0.429 | 0.94697
|  |  |  | -0.286 | 0.94644
|  |  |  | -0.143 | 0.94691
|  |  |  | 0 | 0.94707
|  |  |  | 0.143 | 0.94691
|  |  |  | 0.286 | 0.94644
|  |  |  | 0.429 | 0.94697
|  |  |  | 0.572 | 0.94907
|  |  |  | 0.715 | 0.95086
|  |  |  | 0.858 | 0.95238
|  |  |  | 1.001 | 0.95369
|  |  |  | 1.144 | 0.95484
|  |  |  | 1.287 | 0.95681
|  |  |  | 1.43 | 0.95991
|  |  |  | 1.573 | 0.9628
|  |  |  | 1.716 | 0.96552
|  |  |  | 1.859 | 0.96811
|  |  |  | 2.002 | 0.97058
|  |  |  | 2.145 | 0.97295
|  |  |  | 2.288 | 0.97521
|  |  |  | 2.431 | 0.9773
|  |  |  | 2.574 | 0.97914
|  |  |  | 2.717 | 0.98061
|  |  |  | 2.86 | 0.9815
|  |  |  | 3.003 | 0.98156
|  |  |  | 3.146 | 0.98056
|  |  |  | 3.289 | 0.98134
|  |  |  | 3.432 | 0.98058
|  |  |  | 3.575 | 0.97761

## Strehl Ratio vs. Wavelength
TL10X-2P Strehl Ratio
|  | Wavelength (nm) | Strehl Ratio
|  |  | On Axis | 1° Off Axis | 1.65° Off Axis | 3.25° Off Axis
|  | 400 | 0.9786 | 0.9399 | 0.9005 | 0.7907
|  | 406 | 0.9983 | 0.9659 | 0.9353 | 0.8641
|  | 412 | 0.9933 | 0.9665 | 0.9434 | 0.9078
|  | 418 | 0.9798 | 0.9576 | 0.9403 | 0.9342
Product Raw Data |  | 424 | 0.9653 | 0.9466 | 0.9336 | 0.9506
Thorlabs Super Apochromatic Microscope Objective, 10X Magnification |  | 430 | 0.9523 | 0.9364 | 0.9268 | 0.961
|  | 436 | 0.9415 | 0.928 | 0.9209 | 0.9678
Item # | TL10X-2P | 442 | 0.9329 | 0.9213 | 0.9163 | 0.9723
DISCLAIMER: The data presented here are typical. Slight variations in performance data will occur from lot to lot. Please contact Technical Support with any questions regarding the use or reliability of this data. |  | 448 | 0.9259 | 0.916 | 0.9128 | 0.9753
|  | 454 | 0.92 | 0.9118 | 0.9103 | 0.9772
|  | 460 | 0.915 | 0.9084 | 0.9084 | 0.9783
|  | 466 | 0.9104 | 0.9054 | 0.907 | 0.9791
|  | 472 | 0.9061 | 0.9026 | 0.9059 | 0.9795
|  | 478 | 0.9019 | 0.9 | 0.905 | 0.9799
This data may be used in publications. However, please cite Thorlabs as the source. |  | 484 | 0.8977 | 0.8974 | 0.9042 | 0.9801
|  | 490 | 0.8934 | 0.8947 | 0.9034 | 0.9803
|  | 496 | 0.8891 | 0.892 | 0.9025 | 0.9805
Additional Information: |  | 502 | 0.8848 | 0.8892 | 0.9016 | 0.9808
|  | 508 | 0.8804 | 0.8863 | 0.9007 | 0.981
|  | 514 | 0.876 | 0.8834 | 0.8997 | 0.9812
|  | 520 | 0.8716 | 0.8805 | 0.8986 | 0.9814
|  | 526 | 0.8673 | 0.8776 | 0.8975 | 0.9816
|  | 532 | 0.8632 | 0.8748 | 0.8964 | 0.9817
|  | 538 | 0.8593 | 0.8721 | 0.8954 | 0.9817
|  | 544 | 0.8556 | 0.8696 | 0.8944 | 0.9817
|  | 550 | 0.8522 | 0.8673 | 0.8935 | 0.9816
|  | 556 | 0.8493 | 0.8653 | 0.8928 | 0.9815
|  | 562 | 0.8467 | 0.8636 | 0.8923 | 0.9813
|  | 568 | 0.8446 | 0.8624 | 0.892 | 0.9811
|  | 574 | 0.8431 | 0.8615 | 0.892 | 0.9808
|  | 580 | 0.8421 | 0.861 | 0.8923 | 0.9806
|  | 586 | 0.8416 | 0.8611 | 0.8929 | 0.9803
|  | 592 | 0.8418 | 0.8617 | 0.8939 | 0.98
|  | 598 | 0.8426 | 0.8627 | 0.8952 | 0.9797
|  | 604 | 0.844 | 0.8643 | 0.897 | 0.9793
|  | 610 | 0.846 | 0.8665 | 0.8991 | 0.979
|  | 616 | 0.8487 | 0.8691 | 0.9016 | 0.9786
|  | 622 | 0.8519 | 0.8723 | 0.9045 | 0.9782
|  | 628 | 0.8557 | 0.876 | 0.9077 | 0.9776
|  | 634 | 0.8601 | 0.8801 | 0.9112 | 0.977
|  | 640 | 0.865 | 0.8847 | 0.915 | 0.9762
|  | 646 | 0.8704 | 0.8896 | 0.919 | 0.9752
|  | 652 | 0.8763 | 0.8949 | 0.9233 | 0.974
|  | 658 | 0.8825 | 0.9005 | 0.9276 | 0.9725
|  | 664 | 0.889 | 0.9063 | 0.932 | 0.9706
|  | 670 | 0.8957 | 0.9122 | 0.9364 | 0.9684
|  | 676 | 0.9026 | 0.9183 | 0.9408 | 0.9656
|  | 682 | 0.9096 | 0.9242 | 0.9449 | 0.9623
|  | 688 | 0.9166 | 0.9301 | 0.9488 | 0.9584
|  | 694 | 0.9235 | 0.9358 | 0.9524 | 0.9538
|  | 700 | 0.9301 | 0.9412 | 0.9555 | 0.9484

## Group Delay Dispersion
TL10X-2P Group Delay Dispersion
|  | Wavelength(nm) | GDD (fs2)
|  | 400 | 11421.633
|  | 401 | 11349.8984
|  | 402 | 11279.2201
|  | 403 | 11209.5754
Product Raw Data |  | 404 | 11140.942
Thorlabs Super Apochromatic Microscope Objective, 10X Magnification |  | 405 | 11073.2984
|  | 406 | 11006.6235
Item # | TL10X-2P | 407 | 10940.8968
DISCLAIMER: The data presented here are typical. Slight variations in performance data will occur from lot to lot. Please contact Technical Support with any questions regarding the use or reliability of this data. |  | 408 | 10876.0985
|  | 409 | 10812.209
|  | 410 | 10749.2093
|  | 411 | 10687.0811
|  | 412 | 10625.8062
|  | 413 | 10565.3673
This data may be used in publications. However, please cite Thorlabs as the source. |  | 414 | 10505.7471
|  | 415 | 10446.929
|  | 416 | 10388.8968
Additional Information: |  | 417 | 10331.6345
|  | 418 | 10275.1269
|  | 419 | 10219.3587
|  | 420 | 10164.3153
|  | 421 | 10109.9824
|  | 422 | 10056.346
|  | 423 | 10003.3923
|  | 424 | 9951.1083
|  | 425 | 9899.4807
|  | 426 | 9848.497
|  | 427 | 9798.1448
|  | 428 | 9748.412
|  | 429 | 9699.2868
|  | 430 | 9650.7578
|  | 431 | 9602.8136
|  | 432 | 9555.4434
|  | 433 | 9508.6364
|  | 434 | 9462.3821
|  | 435 | 9416.6704
|  | 436 | 9371.4912
|  | 437 | 9326.8348
|  | 438 | 9282.6918
|  | 439 | 9239.0527
|  | 440 | 9195.9085
|  | 441 | 9153.2504
|  | 442 | 9111.0696
|  | 443 | 9069.3576
|  | 444 | 9028.1062
|  | 445 | 8987.3073
|  | 446 | 8946.9528
|  | 447 | 8907.0352
|  | 448 | 8867.5468
|  | 449 | 8828.4802
|  | 450 | 8789.8281
|  | 451 | 8751.5835
|  | 452 | 8713.7394
|  | 453 | 8676.2891
|  | 454 | 8639.2259
|  | 455 | 8602.5434
|  | 456 | 8566.2352
|  | 457 | 8530.295
|  | 458 | 8494.7168
|  | 459 | 8459.4947
|  | 460 | 8424.6228
|  | 461 | 8390.0954
|  | 462 | 8355.907
|  | 463 | 8322.052
|  | 464 | 8288.5252
|  | 465 | 8255.3212
|  | 466 | 8222.4351
|  | 467 | 8189.8616
|  | 468 | 8157.596
|  | 469 | 8125.6334
|  | 470 | 8093.969
|  | 471 | 8062.5983
|  | 472 | 8031.5167
|  | 473 | 8000.7198
|  | 474 | 7970.2032
|  | 475 | 7939.9627
|  | 476 | 7909.9941
|  | 477 | 7880.2933
|  | 478 | 7850.8562
|  | 479 | 7821.679
|  | 480 | 7792.7578
|  | 481 | 7764.0887
|  | 482 | 7735.6682
|  | 483 | 7707.4925
|  | 484 | 7679.5581
|  | 485 | 7651.8615
|  | 486 | 7624.3993
|  | 487 | 7597.168
|  | 488 | 7570.1645
|  | 489 | 7543.3853
|  | 490 | 7516.8275
|  | 491 | 7490.4878
|  | 492 | 7464.3633
|  | 493 | 7438.4508
|  | 494 | 7412.7474
|  | 495 | 7387.2504
|  | 496 | 7361.9567
|  | 497 | 7336.8637
|  | 498 | 7311.9685
|  | 499 | 7287.2686
|  | 500 | 7262.7613
|  | 501 | 7238.4439
|  | 502 | 7214.314
|  | 503 | 7190.369
|  | 504 | 7166.6065
|  | 505 | 7143.0241
|  | 506 | 7119.6194
|  | 507 | 7096.3901
|  | 508 | 7073.3339
|  | 509 | 7050.4486
|  | 510 | 7027.7319
|  | 511 | 7005.1817
|  | 512 | 6982.7959
|  | 513 | 6960.5724
|  | 514 | 6938.5091
|  | 515 | 6916.604
|  | 516 | 6894.8551
|  | 517 | 6873.2604
|  | 518 | 6851.8181
|  | 519 | 6830.5263
|  | 520 | 6809.383
|  | 521 | 6788.3865
|  | 522 | 6767.535
|  | 523 | 6746.8267
|  | 524 | 6726.2599
|  | 525 | 6705.8328
|  | 526 | 6685.5439
|  | 527 | 6665.3914
|  | 528 | 6645.3737
|  | 529 | 6625.4892
|  | 530 | 6605.7364
|  | 531 | 6586.1137
|  | 532 | 6566.6195
|  | 533 | 6547.2525
|  | 534 | 6528.011
|  | 535 | 6508.8937
|  | 536 | 6489.8991
|  | 537 | 6471.0259
|  | 538 | 6452.2726
|  | 539 | 6433.6378
|  | 540 | 6415.1203
|  | 541 | 6396.7187
|  | 542 | 6378.4318
|  | 543 | 6360.2581
|  | 544 | 6342.1966
|  | 545 | 6324.2458
|  | 546 | 6306.4047
|  | 547 | 6288.672
|  | 548 | 6271.0465
|  | 549 | 6253.5271
|  | 550 | 6236.1126
|  | 551 | 6218.8018
|  | 552 | 6201.5937
|  | 553 | 6184.4872
|  | 554 | 6167.4811
|  | 555 | 6150.5744
|  | 556 | 6133.7661
|  | 557 | 6117.0551
|  | 558 | 6100.4404
|  | 559 | 6083.9209
|  | 560 | 6067.4958
|  | 561 | 6051.164
|  | 562 | 6034.9245
|  | 563 | 6018.7764
|  | 564 | 6002.7188
|  | 565 | 5986.7507
|  | 566 | 5970.8712
|  | 567 | 5955.0794
|  | 568 | 5939.3745
|  | 569 | 5923.7556
|  | 570 | 5908.2218
|  | 571 | 5892.7722
|  | 572 | 5877.406
|  | 573 | 5862.1224
|  | 574 | 5846.9206
|  | 575 | 5831.7998
|  | 576 | 5816.7591
|  | 577 | 5801.7978
|  | 578 | 5786.9152
|  | 579 | 5772.1104
|  | 580 | 5757.3827
|  | 581 | 5742.7314
|  | 582 | 5728.1557
|  | 583 | 5713.6549
|  | 584 | 5699.2283
|  | 585 | 5684.8752
|  | 586 | 5670.595
|  | 587 | 5656.3868
|  | 588 | 5642.2501
|  | 589 | 5628.1841
|  | 590 | 5614.1883
|  | 591 | 5600.2619
|  | 592 | 5586.4043
|  | 593 | 5572.6149
|  | 594 | 5558.8931
|  | 595 | 5545.2382
|  | 596 | 5531.6497
|  | 597 | 5518.1268
|  | 598 | 5504.6691
|  | 599 | 5491.276
|  | 600 | 5477.9468
|  | 601 | 5464.681
|  | 602 | 5451.478
|  | 603 | 5438.3373
|  | 604 | 5425.2583
|  | 605 | 5412.2404
|  | 606 | 5399.2832
|  | 607 | 5386.3861
|  | 608 | 5373.5485
|  | 609 | 5360.77
|  | 610 | 5348.05
|  | 611 | 5335.388
|  | 612 | 5322.7836
|  | 613 | 5310.2362
|  | 614 | 5297.7453
|  | 615 | 5285.3104
|  | 616 | 5272.9312
|  | 617 | 5260.607
|  | 618 | 5248.3375
|  | 619 | 5236.1221
|  | 620 | 5223.9604
|  | 621 | 5211.852
|  | 622 | 5199.7964
|  | 623 | 5187.7932
|  | 624 | 5175.8419
|  | 625 | 5163.9421
|  | 626 | 5152.0934
|  | 627 | 5140.2954
|  | 628 | 5128.5476
|  | 629 | 5116.8496
|  | 630 | 5105.201
|  | 631 | 5093.6015
|  | 632 | 5082.0505
|  | 633 | 5070.5478
|  | 634 | 5059.0929
|  | 635 | 5047.6855
|  | 636 | 5036.3251
|  | 637 | 5025.0114
|  | 638 | 5013.744
|  | 639 | 5002.5226
|  | 640 | 4991.3467
|  | 641 | 4980.216
|  | 642 | 4969.1303
|  | 643 | 4958.089
|  | 644 | 4947.0918
|  | 645 | 4936.1385
|  | 646 | 4925.2286
|  | 647 | 4914.3619
|  | 648 | 4903.5379
|  | 649 | 4892.7564
|  | 650 | 4882.017
|  | 651 | 4871.3195
|  | 652 | 4860.6634
|  | 653 | 4850.0484
|  | 654 | 4839.4744
|  | 655 | 4828.9408
|  | 656 | 4818.4475
|  | 657 | 4807.9941
|  | 658 | 4797.5803
|  | 659 | 4787.2058
|  | 660 | 4776.8703
|  | 661 | 4766.5736
|  | 662 | 4756.3153
|  | 663 | 4746.0951
|  | 664 | 4735.9128
|  | 665 | 4725.7681
|  | 666 | 4715.6607
|  | 667 | 4705.5903
|  | 668 | 4695.5566
|  | 669 | 4685.5595
|  | 670 | 4675.5985
|  | 671 | 4665.6735
|  | 672 | 4655.7841
|  | 673 | 4645.9302
|  | 674 | 4636.1114
|  | 675 | 4626.3276
|  | 676 | 4616.5783
|  | 677 | 4606.8635
|  | 678 | 4597.1828
|  | 679 | 4587.5361
|  | 680 | 4577.9229
|  | 681 | 4568.3432
|  | 682 | 4558.7967
|  | 683 | 4549.2831
|  | 684 | 4539.8022
|  | 685 | 4530.3538
|  | 686 | 4520.9376
|  | 687 | 4511.5534
|  | 688 | 4502.201
|  | 689 | 4492.8802
|  | 690 | 4483.5907
|  | 691 | 4474.3323
|  | 692 | 4465.1048
|  | 693 | 4455.908
|  | 694 | 4446.7416
|  | 695 | 4437.6055
|  | 696 | 4428.4994
|  | 697 | 4419.4232
|  | 698 | 4410.3766
|  | 699 | 4401.3594
|  | 700 | 4392.3715
|  | 701 | 4383.4125
|  | 702 | 4374.4824
|  | 703 | 4365.5808
|  | 704 | 4356.7077
|  | 705 | 4347.8628
|  | 706 | 4339.046
|  | 707 | 4330.257
|  | 708 | 4321.4956
|  | 709 | 4312.7617
|  | 710 | 4304.0551
|  | 711 | 4295.3755
|  | 712 | 4286.7229
|  | 713 | 4278.097
|  | 714 | 4269.4976
|  | 715 | 4260.9246
|  | 716 | 4252.3778
|  | 717 | 4243.857
|  | 718 | 4235.362
|  | 719 | 4226.8928
|  | 720 | 4218.449
|  | 721 | 4210.0305
|  | 722 | 4201.6372
|  | 723 | 4193.2688
|  | 724 | 4184.9253
|  | 725 | 4176.6064
|  | 726 | 4168.3121
|  | 727 | 4160.0421
|  | 728 | 4151.7962
|  | 729 | 4143.5744
|  | 730 | 4135.3764
|  | 731 | 4127.2021
|  | 732 | 4119.0513
|  | 733 | 4110.924
|  | 734 | 4102.8198
|  | 735 | 4094.7388
|  | 736 | 4086.6807
|  | 737 | 4078.6454
|  | 738 | 4070.6327
|  | 739 | 4062.6425
|  | 740 | 4054.6746
|  | 741 | 4046.7289
|  | 742 | 4038.8053
|  | 743 | 4030.9036
|  | 744 | 4023.0237
|  | 745 | 4015.1654
|  | 746 | 4007.3286
|  | 747 | 3999.5131
|  | 748 | 3991.7189
|  | 749 | 3983.9457
|  | 750 | 3976.1935
|  | 751 | 3968.4621
|  | 752 | 3960.7514
|  | 753 | 3953.0612
|  | 754 | 3945.3914
|  | 755 | 3937.7419
|  | 756 | 3930.1126
|  | 757 | 3922.5033
|  | 758 | 3914.9139
|  | 759 | 3907.3443
|  | 760 | 3899.7943
|  | 761 | 3892.2638
|  | 762 | 3884.7528
|  | 763 | 3877.261
|  | 764 | 3869.7883
|  | 765 | 3862.3348
|  | 766 | 3854.9001
|  | 767 | 3847.4842
|  | 768 | 3840.087
|  | 769 | 3832.7084
|  | 770 | 3825.3482
|  | 771 | 3818.0063
|  | 772 | 3810.6826
|  | 773 | 3803.3771
|  | 774 | 3796.0895
|  | 775 | 3788.8198
|  | 776 | 3781.5679
|  | 777 | 3774.3336
|  | 778 | 3767.1169
|  | 779 | 3759.9176
|  | 780 | 3752.7356
|  | 781 | 3745.5709
|  | 782 | 3738.4232
|  | 783 | 3731.2926
|  | 784 | 3724.1788
|  | 785 | 3717.0819
|  | 786 | 3710.0016
|  | 787 | 3702.9379
|  | 788 | 3695.8907
|  | 789 | 3688.8599
|  | 790 | 3681.8454
|  | 791 | 3674.847
|  | 792 | 3667.8648
|  | 793 | 3660.8985
|  | 794 | 3653.948
|  | 795 | 3647.0134
|  | 796 | 3640.0945
|  | 797 | 3633.1911
|  | 798 | 3626.3033
|  | 799 | 3619.4308
|  | 800 | 3612.5737
|  | 801 | 3605.7317
|  | 802 | 3598.9049
|  | 803 | 3592.0931
|  | 804 | 3585.2963
|  | 805 | 3578.5142
|  | 806 | 3571.747
|  | 807 | 3564.9943
|  | 808 | 3558.2563
|  | 809 | 3551.5327
|  | 810 | 3544.8236
|  | 811 | 3538.1287
|  | 812 | 3531.448
|  | 813 | 3524.7815
|  | 814 | 3518.129
|  | 815 | 3511.4905
|  | 816 | 3504.8659
|  | 817 | 3498.255
|  | 818 | 3491.6578
|  | 819 | 3485.0743
|  | 820 | 3478.5043
|  | 821 | 3471.9477
|  | 822 | 3465.4046
|  | 823 | 3458.8747
|  | 824 | 3452.358
|  | 825 | 3445.8544
|  | 826 | 3439.3639
|  | 827 | 3432.8864
|  | 828 | 3426.4217
|  | 829 | 3419.9699
|  | 830 | 3413.5308
|  | 831 | 3407.1044
|  | 832 | 3400.6905
|  | 833 | 3394.2891
|  | 834 | 3387.9002
|  | 835 | 3381.5236
|  | 836 | 3375.1592
|  | 837 | 3368.8071
|  | 838 | 3362.4671
|  | 839 | 3356.1392
|  | 840 | 3349.8232
|  | 841 | 3343.5191
|  | 842 | 3337.2269
|  | 843 | 3330.9464
|  | 844 | 3324.6776
|  | 845 | 3318.4204
|  | 846 | 3312.1748
|  | 847 | 3305.9406
|  | 848 | 3299.7179
|  | 849 | 3293.5065
|  | 850 | 3287.3063
|  | 851 | 3281.1173
|  | 852 | 3274.9395
|  | 853 | 3268.7727
|  | 854 | 3262.617
|  | 855 | 3256.4721
|  | 856 | 3250.3381
|  | 857 | 3244.2149
|  | 858 | 3238.1024
|  | 859 | 3232.0006
|  | 860 | 3225.9094
|  | 861 | 3219.8287
|  | 862 | 3213.7584
|  | 863 | 3207.6986
|  | 864 | 3201.6491
|  | 865 | 3195.6098
|  | 866 | 3189.5808
|  | 867 | 3183.5619
|  | 868 | 3177.5532
|  | 869 | 3171.5544
|  | 870 | 3165.5656
|  | 871 | 3159.5867
|  | 872 | 3153.6176
|  | 873 | 3147.6583
|  | 874 | 3141.7087
|  | 875 | 3135.7688
|  | 876 | 3129.8384
|  | 877 | 3123.9176
|  | 878 | 3118.0063
|  | 879 | 3112.1044
|  | 880 | 3106.2118
|  | 881 | 3100.3286
|  | 882 | 3094.4546
|  | 883 | 3088.5898
|  | 884 | 3082.7341
|  | 885 | 3076.8875
|  | 886 | 3071.0499
|  | 887 | 3065.2212
|  | 888 | 3059.4015
|  | 889 | 3053.5906
|  | 890 | 3047.7885
|  | 891 | 3041.9951
|  | 892 | 3036.2104
|  | 893 | 3030.4344
|  | 894 | 3024.6669
|  | 895 | 3018.9079
|  | 896 | 3013.1574
|  | 897 | 3007.4153
|  | 898 | 3001.6815
|  | 899 | 2995.9561
|  | 900 | 2990.2389
|  | 901 | 2984.5299
|  | 902 | 2978.8291
|  | 903 | 2973.1364
|  | 904 | 2967.4517
|  | 905 | 2961.7749
|  | 906 | 2956.1062
|  | 907 | 2950.4453
|  | 908 | 2944.7923
|  | 909 | 2939.147
|  | 910 | 2933.5095
|  | 911 | 2927.8797
|  | 912 | 2922.2575
|  | 913 | 2916.643
|  | 914 | 2911.0359
|  | 915 | 2905.4364
|  | 916 | 2899.8443
|  | 917 | 2894.2596
|  | 918 | 2888.6822
|  | 919 | 2883.1122
|  | 920 | 2877.5494
|  | 921 | 2871.9938
|  | 922 | 2866.4453
|  | 923 | 2860.904
|  | 924 | 2855.3697
|  | 925 | 2849.8425
|  | 926 | 2844.3222
|  | 927 | 2838.8089
|  | 928 | 2833.3024
|  | 929 | 2827.8028
|  | 930 | 2822.31
|  | 931 | 2816.8239
|  | 932 | 2811.3445
|  | 933 | 2805.8717
|  | 934 | 2800.4056
|  | 935 | 2794.946
|  | 936 | 2789.493
|  | 937 | 2784.0464
|  | 938 | 2778.6063
|  | 939 | 2773.1726
|  | 940 | 2767.7452
|  | 941 | 2762.3241
|  | 942 | 2756.9093
|  | 943 | 2751.5007
|  | 944 | 2746.0982
|  | 945 | 2740.702
|  | 946 | 2735.3118
|  | 947 | 2729.9276
|  | 948 | 2724.5495
|  | 949 | 2719.1773
|  | 950 | 2713.8111
|  | 951 | 2708.4507
|  | 952 | 2703.0962
|  | 953 | 2697.7475
|  | 954 | 2692.4045
|  | 955 | 2687.0673
|  | 956 | 2681.7358
|  | 957 | 2676.4099
|  | 958 | 2671.0896
|  | 959 | 2665.7749
|  | 960 | 2660.4657
|  | 961 | 2655.1619
|  | 962 | 2649.8637
|  | 963 | 2644.5708
|  | 964 | 2639.2833
|  | 965 | 2634.0011
|  | 966 | 2628.7242
|  | 967 | 2623.4526
|  | 968 | 2618.1862
|  | 969 | 2612.925
|  | 970 | 2607.6689
|  | 971 | 2602.4179
|  | 972 | 2597.1719
|  | 973 | 2591.931
|  | 974 | 2586.6951
|  | 975 | 2581.4641
|  | 976 | 2576.2381
|  | 977 | 2571.0169
|  | 978 | 2565.8006
|  | 979 | 2560.5891
|  | 980 | 2555.3823
|  | 981 | 2550.1803
|  | 982 | 2544.983
|  | 983 | 2539.7903
|  | 984 | 2534.6023
|  | 985 | 2529.4189
|  | 986 | 2524.24
|  | 987 | 2519.0656
|  | 988 | 2513.8958
|  | 989 | 2508.7303
|  | 990 | 2503.5693
|  | 991 | 2498.4127
|  | 992 | 2493.2604
|  | 993 | 2488.1125
|  | 994 | 2482.9688
|  | 995 | 2477.8294
|  | 996 | 2472.6942
|  | 997 | 2467.5631
|  | 998 | 2462.4362
|  | 999 | 2457.3134
|  | 1000 | 2452.1947
|  | 1001 | 2447.0801
|  | 1002 | 2441.9694
|  | 1003 | 2436.8627
|  | 1004 | 2431.76
|  | 1005 | 2426.6612
|  | 1006 | 2421.5663
|  | 1007 | 2416.4752
|  | 1008 | 2411.3879
|  | 1009 | 2406.3044
|  | 1010 | 2401.2247
|  | 1011 | 2396.1487
|  | 1012 | 2391.0764
|  | 1013 | 2386.0077
|  | 1014 | 2380.9427
|  | 1015 | 2375.8812
|  | 1016 | 2370.8234
|  | 1017 | 2365.769
|  | 1018 | 2360.7182
|  | 1019 | 2355.6708
|  | 1020 | 2350.6269
|  | 1021 | 2345.5863
|  | 1022 | 2340.5492
|  | 1023 | 2335.5154
|  | 1024 | 2330.4849
|  | 1025 | 2325.4577
|  | 1026 | 2320.4338
|  | 1027 | 2315.4131
|  | 1028 | 2310.3956
|  | 1029 | 2305.3812
|  | 1030 | 2300.37
|  | 1031 | 2295.362
|  | 1032 | 2290.3569
|  | 1033 | 2285.355
|  | 1034 | 2280.356
|  | 1035 | 2275.3601
|  | 1036 | 2270.3671
|  | 1037 | 2265.3771
|  | 1038 | 2260.3899
|  | 1039 | 2255.4056
|  | 1040 | 2250.4242
|  | 1041 | 2245.4456
|  | 1042 | 2240.4698
|  | 1043 | 2235.4968
|  | 1044 | 2230.5265
|  | 1045 | 2225.5589
|  | 1046 | 2220.594
|  | 1047 | 2215.6317
|  | 1048 | 2210.672
|  | 1049 | 2205.715
|  | 1050 | 2200.7605
|  | 1051 | 2195.8086
|  | 1052 | 2190.8592
|  | 1053 | 2185.9122
|  | 1054 | 2180.9678
|  | 1055 | 2176.0257
|  | 1056 | 2171.0861
|  | 1057 | 2166.1488
|  | 1058 | 2161.2139
|  | 1059 | 2156.2813
|  | 1060 | 2151.3511
|  | 1061 | 2146.4231
|  | 1062 | 2141.4973
|  | 1063 | 2136.5738
|  | 1064 | 2131.6524
|  | 1065 | 2126.7332
|  | 1066 | 2121.8162
|  | 1067 | 2116.9013
|  | 1068 | 2111.9885
|  | 1069 | 2107.0777
|  | 1070 | 2102.169
|  | 1071 | 2097.2623
|  | 1072 | 2092.3575
|  | 1073 | 2087.4548
|  | 1074 | 2082.5539
|  | 1075 | 2077.655
|  | 1076 | 2072.758
|  | 1077 | 2067.8628
|  | 1078 | 2062.9695
|  | 1079 | 2058.078
|  | 1080 | 2053.1883
|  | 1081 | 2048.3003
|  | 1082 | 2043.414
|  | 1083 | 2038.5295
|  | 1084 | 2033.6467
|  | 1085 | 2028.7655
|  | 1086 | 2023.886
|  | 1087 | 2019.0081
|  | 1088 | 2014.1318
|  | 1089 | 2009.257
|  | 1090 | 2004.3838
|  | 1091 | 1999.5121
|  | 1092 | 1994.6419
|  | 1093 | 1989.7732
|  | 1094 | 1984.9059
|  | 1095 | 1980.0401
|  | 1096 | 1975.1756
|  | 1097 | 1970.3125
|  | 1098 | 1965.4508
|  | 1099 | 1960.5904
|  | 1100 | 1955.7313
|  | 1101 | 1950.8735
|  | 1102 | 1946.017
|  | 1103 | 1941.1617
|  | 1104 | 1936.3076
|  | 1105 | 1931.4547
|  | 1106 | 1926.603
|  | 1107 | 1921.7524
|  | 1108 | 1916.9029
|  | 1109 | 1912.0545
|  | 1110 | 1907.2073
|  | 1111 | 1902.361
|  | 1112 | 1897.5158
|  | 1113 | 1892.6717
|  | 1114 | 1887.8285
|  | 1115 | 1882.9862
|  | 1116 | 1878.145
|  | 1117 | 1873.3046
|  | 1118 | 1868.4651
|  | 1119 | 1863.6265
|  | 1120 | 1858.7888
|  | 1121 | 1853.9519
|  | 1122 | 1849.1158
|  | 1123 | 1844.2805
|  | 1124 | 1839.446
|  | 1125 | 1834.6122
|  | 1126 | 1829.7792
|  | 1127 | 1824.9468
|  | 1128 | 1820.1152
|  | 1129 | 1815.2842
|  | 1130 | 1810.4538
|  | 1131 | 1805.6241
|  | 1132 | 1800.7949
|  | 1133 | 1795.9663
|  | 1134 | 1791.1383
|  | 1135 | 1786.3108
|  | 1136 | 1781.4839
|  | 1137 | 1776.6574
|  | 1138 | 1771.8314
|  | 1139 | 1767.0059
|  | 1140 | 1762.1807
|  | 1141 | 1757.356
|  | 1142 | 1752.5317
|  | 1143 | 1747.7078
|  | 1144 | 1742.8842
|  | 1145 | 1738.0609
|  | 1146 | 1733.2379
|  | 1147 | 1728.4152
|  | 1148 | 1723.5928
|  | 1149 | 1718.7707
|  | 1150 | 1713.9487
|  | 1151 | 1709.127
|  | 1152 | 1704.3054
|  | 1153 | 1699.4841
|  | 1154 | 1694.6628
|  | 1155 | 1689.8417
|  | 1156 | 1685.0207
|  | 1157 | 1680.1998
|  | 1158 | 1675.379
|  | 1159 | 1670.5582
|  | 1160 | 1665.7374
|  | 1161 | 1660.9167
|  | 1162 | 1656.0959
|  | 1163 | 1651.2751
|  | 1164 | 1646.4542
|  | 1165 | 1641.6333
|  | 1166 | 1636.8123
|  | 1167 | 1631.9912
|  | 1168 | 1627.17
|  | 1169 | 1622.3486
|  | 1170 | 1617.527
|  | 1171 | 1612.7053
|  | 1172 | 1607.8833
|  | 1173 | 1603.0612
|  | 1174 | 1598.2388
|  | 1175 | 1593.4161
|  | 1176 | 1588.5932
|  | 1177 | 1583.7699
|  | 1178 | 1578.9464
|  | 1179 | 1574.1225
|  | 1180 | 1569.2982
|  | 1181 | 1564.4736
|  | 1182 | 1559.6486
|  | 1183 | 1554.8232
|  | 1184 | 1549.9973
|  | 1185 | 1545.1711
|  | 1186 | 1540.3443
|  | 1187 | 1535.5171
|  | 1188 | 1530.6893
|  | 1189 | 1525.8611
|  | 1190 | 1521.0323
|  | 1191 | 1516.2029
|  | 1192 | 1511.373
|  | 1193 | 1506.5425
|  | 1194 | 1501.7113
|  | 1195 | 1496.8796
|  | 1196 | 1492.0472
|  | 1197 | 1487.2141
|  | 1198 | 1482.3804
|  | 1199 | 1477.5459
|  | 1200 | 1472.7107
|  | 1201 | 1467.8748
|  | 1202 | 1463.0382
|  | 1203 | 1458.2008
|  | 1204 | 1453.3625
|  | 1205 | 1448.5235
|  | 1206 | 1443.6837
|  | 1207 | 1438.843
|  | 1208 | 1434.0015
|  | 1209 | 1429.159
|  | 1210 | 1424.3157
|  | 1211 | 1419.4715
|  | 1212 | 1414.6264
|  | 1213 | 1409.7803
|  | 1214 | 1404.9332
|  | 1215 | 1400.0852
|  | 1216 | 1395.2361
|  | 1217 | 1390.3861
|  | 1218 | 1385.535
|  | 1219 | 1380.6829
|  | 1220 | 1375.8297
|  | 1221 | 1370.9754
|  | 1222 | 1366.1201
|  | 1223 | 1361.2636
|  | 1224 | 1356.406
|  | 1225 | 1351.5472
|  | 1226 | 1346.6873
|  | 1227 | 1341.8261
|  | 1228 | 1336.9638
|  | 1229 | 1332.1003
|  | 1230 | 1327.2355
|  | 1231 | 1322.3695
|  | 1232 | 1317.5022
|  | 1233 | 1312.6337
|  | 1234 | 1307.7638
|  | 1235 | 1302.8926
|  | 1236 | 1298.0201
|  | 1237 | 1293.1462
|  | 1238 | 1288.271
|  | 1239 | 1283.3944
|  | 1240 | 1278.5164
|  | 1241 | 1273.637
|  | 1242 | 1268.7561
|  | 1243 | 1263.8738
|  | 1244 | 1258.9901
|  | 1245 | 1254.1048
|  | 1246 | 1249.2181
|  | 1247 | 1244.3299
|  | 1248 | 1239.4401
|  | 1249 | 1234.5488
|  | 1250 | 1229.6559
|  | 1251 | 1224.7614
|  | 1252 | 1219.8654
|  | 1253 | 1214.9677
|  | 1254 | 1210.0685
|  | 1255 | 1205.1676
|  | 1256 | 1200.265
|  | 1257 | 1195.3608
|  | 1258 | 1190.4548
|  | 1259 | 1185.5472
|  | 1260 | 1180.6379
|  | 1261 | 1175.7268
|  | 1262 | 1170.814
|  | 1263 | 1165.8994
|  | 1264 | 1160.983
|  | 1265 | 1156.0648
|  | 1266 | 1151.1448
|  | 1267 | 1146.223
|  | 1268 | 1141.2994
|  | 1269 | 1136.3739
|  | 1270 | 1131.4465
|  | 1271 | 1126.5172
|  | 1272 | 1121.5861
|  | 1273 | 1116.653
|  | 1274 | 1111.7179
|  | 1275 | 1106.781
|  | 1276 | 1101.842
|  | 1277 | 1096.9011
|  | 1278 | 1091.9582
|  | 1279 | 1087.0132
|  | 1280 | 1082.0663
|  | 1281 | 1077.1173
|  | 1282 | 1072.1662
|  | 1283 | 1067.2131
|  | 1284 | 1062.2579
|  | 1285 | 1057.3006
|  | 1286 | 1052.3411
|  | 1287 | 1047.3796
|  | 1288 | 1042.4158
|  | 1289 | 1037.45
|  | 1290 | 1032.4819
|  | 1291 | 1027.5117
|  | 1292 | 1022.5392
|  | 1293 | 1017.5646
|  | 1294 | 1012.5876
|  | 1295 | 1007.6085
|  | 1296 | 1002.6271
|  | 1297 | 997.6434
|  | 1298 | 992.6574
|  | 1299 | 987.6691
|  | 1300 | 982.6784
|  | 1301 | 977.6855
|  | 1302 | 972.6901
|  | 1303 | 967.6924
|  | 1304 | 962.6924
|  | 1305 | 957.6899
|  | 1306 | 952.685
|  | 1307 | 947.6777
|  | 1308 | 942.6679
|  | 1309 | 937.6557
|  | 1310 | 932.641
|  | 1311 | 927.6239
|  | 1312 | 922.6042
|  | 1313 | 917.582
|  | 1314 | 912.5573
|  | 1315 | 907.5301
|  | 1316 | 902.5003
|  | 1317 | 897.4679
|  | 1318 | 892.433
|  | 1319 | 887.3954
|  | 1320 | 882.3552
|  | 1321 | 877.3124
|  | 1322 | 872.267
|  | 1323 | 867.2189
|  | 1324 | 862.1681
|  | 1325 | 857.1147
|  | 1326 | 852.0585
|  | 1327 | 846.9997
|  | 1328 | 841.9381
|  | 1329 | 836.8738
|  | 1330 | 831.8067
|  | 1331 | 826.7368
|  | 1332 | 821.6642
|  | 1333 | 816.5887
|  | 1334 | 811.5105
|  | 1335 | 806.4294
|  | 1336 | 801.3455
|  | 1337 | 796.2588
|  | 1338 | 791.1691
|  | 1339 | 786.0766
|  | 1340 | 780.9812
|  | 1341 | 775.8829
|  | 1342 | 770.7817
|  | 1343 | 765.6775
|  | 1344 | 760.5704
|  | 1345 | 755.4603
|  | 1346 | 750.3473
|  | 1347 | 745.2312
|  | 1348 | 740.1122
|  | 1349 | 734.9901
|  | 1350 | 729.865
|  | 1351 | 724.7369
|  | 1352 | 719.6057
|  | 1353 | 714.4714
|  | 1354 | 709.334
|  | 1355 | 704.1935
|  | 1356 | 699.05
|  | 1357 | 693.9033
|  | 1358 | 688.7534
|  | 1359 | 683.6004
|  | 1360 | 678.4442
|  | 1361 | 673.2849
|  | 1362 | 668.1223
|  | 1363 | 662.9566
|  | 1364 | 657.7876
|  | 1365 | 652.6154
|  | 1366 | 647.4399
|  | 1367 | 642.2612
|  | 1368 | 637.0792
|  | 1369 | 631.8939
|  | 1370 | 626.7053
|  | 1371 | 621.5134
|  | 1372 | 616.3182
|  | 1373 | 611.1196
|  | 1374 | 605.9177
|  | 1375 | 600.7124
|  | 1376 | 595.5037
|  | 1377 | 590.2916
|  | 1378 | 585.0762
|  | 1379 | 579.8572
|  | 1380 | 574.6349
|  | 1381 | 569.4091
|  | 1382 | 564.1799
|  | 1383 | 558.9471
|  | 1384 | 553.7109
|  | 1385 | 548.4712
|  | 1386 | 543.228
|  | 1387 | 537.9812
|  | 1388 | 532.7309
|  | 1389 | 527.4771
|  | 1390 | 522.2197
|  | 1391 | 516.9587
|  | 1392 | 511.6941
|  | 1393 | 506.4259
|  | 1394 | 501.1541
|  | 1395 | 495.8786
|  | 1396 | 490.5995
|  | 1397 | 485.3168
|  | 1398 | 480.0304
|  | 1399 | 474.7403
|  | 1400 | 469.4465